## Legend+Explanations
Explanation
Name of Excel sheet | Contents
"Report" | The first section of page 1 has to be completed before beginning laboratory work! The values for species, target sequence/haploid genome copies, %GMO, haploid genome copy number, DNA concentration, PCR volume and template volume are needed for further calculations (e.g. dilution series, dilution of samples with GM level 1 and GM level 2, preparation of the reaction mix).
| Estimation of LODabs and LOQabs: A dilution series with dilution levels A-I (e.g. 2.500, 500, 100, 60, 40, 20, 10, 5 and 1 GM target DNA copies) is prepared. Dilution levels D-I are measured in 10 replicates. | Plate A
Determination of the dynamic range, R² coefficient and amplification efficiency: 2 calibration curves, 5 calibration points with 3 PCR replicates each (according to example 1 in table 1 in the guidance document from the ENGL "Verification of analytical methods for GMO testing when implementing interlaboratory validated methods")
Plate B + C
Determination of trueness, precision and RSDr : 2 DNA extraction replicates per GM level, 4 PCR replicates per extraction/plate, 2 plates, resulting in 16 test results and 4 GM-estimations per GM level (according to example 2 in table 1 in the guidance document from the ENGL "Verification of analytical methods for GMO testing when implementing interlaboratory validated methods")
| Additionally, new materials can be tested for specificity. | Plate C
The measured Cq-values are entered on page 1 for specificity, on page 2 for plate A (LOD and LOQ) and on page 3 for plates B and C (dynamic range, R² coefficient, amplification efficiency, trueness, precision and RSDr).
The second section of page 1 provides an overview of the determined data. The evaluation of the acceptance criteria is highlighted (fulfilled or not fulfilled).
"Dilution series A" | Preparation of dilution series for plate A
"Dilution series B" | Preparation of dilution series for plate B
"Dilution series C" | Preparation of dilution series for plate C
"Dilution of samples" | Dilution of samples with GM level 1 and GM level 2 and dilution of samples for specificity test
"Plate layout A "LOD_ LOQ"" | Plate A: Plate setup - loading order, cycling programme, preparation of working solutions of primers and probes, preparation of reaction mix
"Plate layout B "Trueness"" | Plate B: Plate setup - loading order, cycling programme, preparation of working solutions of primers and probes (reference and GM target system), preparation of reaction mix (reference and GM target system)
"Plate layout C "Trueness"" | Plate C: Plate setup - loading order, cycling programme, preparation of working solutions of primers and probes (reference and GM target system), preparation of reaction mix (reference and GM target system)
Legend
: This field must be filled out
: This field should be filled out
: The content of this field can be adapted
All other fields are write-protected!
AS-No: | : Internal number of the analysis
S/N | : Serial number (internal material number)
Name | : Name of the material
Sample No | : Sample number (internal number/name of the DNA extraction replicate)
EBC | : Extraction blank control
Primer No | : Internal number of primers and probes
GM level 1 | : Close to the level set in legislation
GM level 2 | : Close to LOQ
PC | : Peformance criteria
}

## Report
Verification report of methods validated in an interlaboratory study for the quantification with real-time PCR of: | [34 empty] | Calculation of values of standard curve
[10 empty] | DAS-81419-2 |  | successful |  |  | Determination of LOD / LOQ | [19 empty] | Log hpl. genom equiv. | Cq |  |  | Values of standard curve | [12 empty] | Trueness GM level 1: | [5 empty] | %GMO: |  |  | 1%
[35 empty] | 3.3979 | 27.24
[14 empty] | Dilution level | GM target DNA copy number [target value] | Cq-values of PCR replicates of Plate A | [10 empty] | Curve parameters | [5 empty] | PC + 10/10 + previous level | LOD | 3.3979 | 27.21 |  |  | reference gene |  |  |  |  | GM target | [5 empty] | Extraction 1 |  |  |  | Sample No. | D150164
| Date | 2015-06-10 |  |  | Species | soybean |  |  | AS-No | 150086 | [16 empty] | intercept [2500 - x] | slope [2500 - x] | R² coefficient [2500 - x] | [5 empty] | 3.3979 | 27.16 | Log haploid genome equivalents | Dilution | Haploid genome copy number [target value] | Pate B |  | Plate C |  | Haploid genome copy number [target value] | Plate B |  | Plate C |  | Log haploid Genome equivalents | Plate B | [7 empty] | Plate C
| Operator | Müller |  |  | Target sequence/hapl. genome copies | 1 |  |  | Reference gene | Lectin | [19 empty] | PC- calibration curve | 10/10 | PC+10/10 |  |  | 2.699 | 29.48 |  |  |  | Haploid genome copy number [actual value] | Cq-value | Haploid genome copy number [actual value] | Cq-value |  | Haploid genome copy number [actual value] | Cq-value | Haploid genome copy number [actual value] | Cq-value |  | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2 |  | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2
Responsible scientist |  | Muster | [6 empty] | Event | DAS-81419-2 | [24 empty] | 2.699 | 29.45 | [13 empty] | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies
[35 empty] | 2.699 | 29.49 |  |  |  |  | Ct value |  | Ct value |  |  | Ct value |  | Ct value |  | 24.24 | 42079.7201 |  |  | 31.17 | 413.9554 |  |  | 24.26 | 42865.0106 |  |  | 30.85 | 454.4972
|  |  | Standard |  |  | GM level 1.1 (Extract. 1) |  | GM level 1.2 (Extract. 2) |  | GM level 2.1 (Extract. 1) |  | GM level 2.2 (Extract. 2) |  | A | 2500 | 27.24 | 27.21 | 27.16 | [7 empty] | 2500 | [8 empty] | 2 | 32.17 | 4.9031 | Level A | 80000 | 77602.1734 | 23.33 | 83465.7224 | 23.24 | 8000 | 7131.5483 | 26.89 | 7664.9096 | 26.61 | 3.9031 | 24.32 | 39875.4611 |  |  | 31.22 | 400.4162 |  |  | 24.28 | 42308.5677 |  |  | 30.91 | 436.6851
| S/N | #311/1 |  |  | S/N | #310/1 |  | #310/1 |  | #309/1 |  | #309/1 |  | B | 500 | 29.48 | 29.45 | 29.49 | [7 empty] | 500 | [8 empty] | 2 | 32.11 | 4.9031 |  | 80000 | 81891.9116 | 23.25 | 77172.1293 | 23.36 | 8000 | 7227.0424 | 26.87 | 7767.7397 | 26.59 | 3.9031 | 24.33 | 39608.1745 |  |  | 31.08 | 439.49 |  |  | 24.31 | 41487.4178 |  |  | 31.08 | 389.9183
| Name | DAS-81419-2 10% |  |  | Name | DAS-81419-2 1,0% |  |  |  | DAS-81419-2 0,1% |  |  |  | C | 100 | 32.17 | 32.11 | 31.89 | [7 empty] | 100 | [8 empty] | 2 | 31.89 | 4.9031 |  | 80000 | 79183.8395 | 23.3 | 72765.3677 | 23.45 | 8000 | 7275.2679 | 26.86 | 7413.7508 | 26.66 | 3.9031 | 24.42 | 37281.7948 |  |  | 31.32 | 374.6517 |  |  | 24.31 | 41487.4178 |  |  | 31.1 | 384.7565
| % GMO | 10% |  |  | % GMO | 1% |  |  |  | 0.1% |  |  |  | D | 60 | 33.03 | 32.65 | 32.91 | 32.47 | 32.6 | 32.37 | 32.9 | 32.88 | 32.61 | 33.11 | 60 | 38.9332 | -3.4669 | 0.99782 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.7782 | 32.753 | 4.2041 | Level B | 16000 | 16411.4956 | 25.64 | 16406.6132 | 25.73 | 1600 | 1729.6454 | 29.02 | 1817.3722 | 28.77 | 3.2041 | % GMO1 |  |  | 1.028% |  | sd1: | 0.00084869 |  | % GMO3 |  |  | 0.991% |  | sd3: | 0.00083534
|  | Hapl. genome copy number |  |  |  | Hapl. genome copy number |  |  |  | Hapl. genome copy number |  |  |  | E | 40 | 33.47 | 33.28 | 33.45 | 33.7 | 33.31 | 33.68 | 33.47 | 33.28 | 33.99 | 33.63 | 40 | 39.0198 | -3.496 | 0.998 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.6021 | 33.526 | 4.2041 |  | 16000 | 17202.6121 | 25.57 | 17286.9107 | 25.65 | 1600 | 1618.3527 | 29.12 | 1655.5052 | 28.91 | 3.2041
| Level A | 80000 |  |  | Extraction 1 | 50000 |  |  |  | 50000 |  |  |  | F | 20 | 34.53 | 34.37 | 34.67 | 34.88 | 34.75 | 33.96 | 34.23 | 34.3 | 34.43 | 34.35 | 20 | 39.0068 | -3.4915 | 0.99845 | TRUE | TRUE | TRUE | TRUE | TRUE | 1.301 | 34.447 | 4.2041 |  | 16000 | 15762.4172 | 25.7 | 16193.6343 | 25.75 | 1600 | 1673.0739 | 29.07 | 1622.7406 | 28.94 | 3.2041 | Extraction 2 |  |  |  | Sample No. | D150168
| Level B | 16000 |  |  | Extraction 2 | 50000 |  |  |  | 50000 |  |  |  | G | 10 | 35.3 | 35.44 | 35.55 | 35.62 | 35.56 | 35.66 | 36.42 | 36.32 | 36.28 | 36.9 | 10 | 39.1869 | -3.5566 | 0.99778 | TRUE | TRUE | TRUE | TRUE | TRUE | 1 | 35.905 | 3.6021 | Level C | 4000 | 3813.4223 | 27.81 | 3948.9916 | 27.91 | 400 | 436.5768 | 31.09 | 428.0425 | 30.94 | 2.6021 | Plate B | [7 empty] | Plate C
| Level C | 4000 |  |  |  | Operator |  | Date of extraction |  | Operator |  | Date of extraction |  | H | 5 | 36.84 | 36.1 | 36.38 | 39.99 | 37.39 | 35.96 | 36.7 | 38.67 | 36.08 | 36.57 | 5 | 39.3307 | -3.6103 | 0.99763 | FALSE | TRUE | FALSE | FALSE | TRUE | 0.69897 | 37.068 | 3.6021 |  | 4000 | 3865.064 | 27.79 | 3872.3472 | 27.94 | 400 | 411.2114 | 31.18 | 414.0167 | 30.99 | 2.6021 | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2 |  | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2
| Level D | 800 |  |  | Extraction 1 | Müller |  | 2015-06-03 |  | Müller |  | 2015-06-03 |  | I | 1 |  | 37.81 | 38.79 |  | 38.42 | 37.41 |  |  | 39.84 | 38.74 | 1 | serial dilution plausible | [9 empty] | 3.6021 |  | 4000 | 4024.2229 | 27.73 | 4133.7799 | 27.84 | 400 | 463.5068 | 31 | 422.3761 | 30.96 | 2.6021 | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies
| Level E | 200 |  |  | Extraction 2 | Müller |  | 2015-06-05 |  | Müller |  | 2015-06-05 | [20 empty] | LOQ |  |  |  | 2.9031 | Level D | 800 | 769.3836 | 30.19 | 791.6057 | 30.37 | 80 | 85.5865 | 33.54 | 71.7704 | 33.62 | 1.9031 | 24.01 | 49119.4308 |  |  | 30.96 | 476.003 |  |  | 24.18 | 45164.9344 |  |  | 30.71 | 498.9357
|  | Sample No. | Conc. (ng/µl) | Sample-No EBC | Extract. Method: | AV111_10_01 Fast ID Genomic | [17 empty] | LODabs |  | 5 | [10 empty] | 2.9031 |  | 800 | 785.065 | 30.16 | 807.2738 | 30.34 | 80 | 90.8659 | 33.45 | 70.8203 | 33.64 | 1.9031 | 24.15 | 44705.4898 |  |  | 31.04 | 451.3387 |  |  | 24.13 | 46664.6394 |  |  | 30.76 | 482.5869
Standard of Plate A |  | D150150 | 123.2 | D150153 |  | Sample No. |  | Conc. (ng/µl) | Sample No. EBC | Sample No. |  | Conc. (ng/µl) | Sample No. EBC | [17 empty] | <25% and 10/10 | <25% and 10/10 PC+10/10 |  |  |  |  | 2.9031 |  | 800 | 828.4623 | 30.08 | 839.5463 | 30.28 | 80 | 93.9383 | 33.4 | 78.2646 | 33.49 | 1.9031 | 24.18 | 43812.5152 |  |  | 30.96 | 476.003 |  |  | 24.11 | 47278.3734 |  |  | 30.88 | 445.5021
Standard of Plate B |  | D150150 | 123.2 | D150153 | Extraction 1 | D150164 |  | 106.9 | D150166 | D150160 |  | 110.2 | D150166 | [23 empty] | 2.301 | Level E | 200 | 210.0928 | 32.12 | 177.3233 | 32.66 | 20 | 17.578 | 35.92 | 19.968 | 35.54 | 1.301 | 24.2 | 43227.1302 |  |  | 31.02 | 457.3823 |  |  | 24.13 | 46664.6394 |  |  | 30.8 | 469.8944
Standard of Plate C |  | D150150 | 123.2 | D150153 | Extraction 2 | D150168 |  | 142.6 | D150169 | D150167 |  | 132.2 | D150169 | [23 empty] | 2.301 |  | 200 | 196.4274 | 32.22 | 202.0747 | 32.46 | 20 | 17.6953 | 35.91 | 19.4428 | 35.58 | 1.301 | % GMO2 |  |  | 1.032% |  | sd2 | 0.00067009 |  | % GMO4 |  |  | 1.021% |  | sd4 | 0.00052413
[37 empty] | 2.301 |  | 200 | 200.431 | 32.19 | 210.1531 | 32.4 | 20 | 16.5567 | 36.01 | 22.0669 | 35.39 | 1.301
| Instrument | Mx3005 | [13 empty] | GM target DNA copy number [actual value] | [9 empty] | RSD |  |  |  | <25% | [19 empty] | Mean GMO %: |  |  | 1.018% |  |  | sdGM | 0.00073175 |  |  | RSDr | 7.187%
| Ref. gene | Lectin | [5 empty] | Event: | DAS-81419-2 | [5 empty] | 2500 | 2285.9197 | 2330.7512 | 2407.4327 | [11 empty] | TRUE | TRUE | TRUE | TRUE | [5 empty] | Efficiency |  | 95.92% |  | 92.19% |  |  | 94.46% |  | 94.71%
| PCR reagent kit | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | [6 empty] | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | [5 empty] | 500 | 536.1073 | 546.6214 | 532.6477 | [11 empty] | TRUE | TRUE | TRUE | TRUE | [5 empty] | Slope |  | -3.4236 |  | -3.5245 |  |  | -3.4621 |  | -3.4557
[15 empty] | 100 | 93.9546 | 97.6761 | 112.6275 | [11 empty] | TRUE | TRUE | TRUE | TRUE | [5 empty] | Intercept |  | 40.071 |  | 40.5857 |  |  | 40.2303 |  | 40.0335 |  |  |  | Trueness GM level 2: | [5 empty] | %GMO: |  |  | 0.1%
|  | Primer F | Primer R | Probe |  |  |  |  | Primer F | Primer R | Probe |  |  |  | 60 | 53.8414 | 68.8587 | 58.1911 | 77.3693 | 71.1242 | 82.544 | 58.569 | 59.3323 | 70.6652 | 51.1238 | 16.03% |  |  |  | FALSE | FALSE | FALSE | FALSE | [5 empty] | R2 |  | 0.99975 |  | 0.9994 |  |  | 0.99726 |  | 0.99893
| Primer No. | R010 | R011 | R012B |  |  |  | Primer No. | R363 | R364 | R365 |  |  |  | 40 | 40.4952 | 45.7958 | 41.023 | 34.8928 | 44.9149 | 35.3475 | 40.4952 | 45.7958 | 28.92 | 36.5105 | 13.95% |  |  |  | FALSE | FALSE | FALSE | FALSE | [16 empty] | Extraction 1 |  |  |  | Sample No. | D150160
PCR Vol. (µl) |  | 25 |  |  |  |  | PCR Vol. (µl) |  | 25 | [5 empty] | 20 | 20.3877 | 22.6129 | 18.6211 | 16.254 | 17.6812 | 29.4872 | 24.7582 | 23.6612 | 21.7513 | 22.9076 | 17.63% |  |  |  | Dynamic Range: |  | 2500 | - | 10 | [15 empty] | Plate B | [7 empty] | Plate C
[6 empty] | Template Vol. (µl) | [8 empty] | 10 | 12.3843 | 11.3112 | 10.5337 | 10.067 | 10.4657 | 9.8096 | 5.9975 | 6.3986 | 6.5665 | 4.3955 | 30.7% |  |  |  | Slope: |  | -3.5566 | [17 empty] | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2 |  | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2
[6 empty] | 5 | [8 empty] | 5 | 4.5696 | 7.378 | 6.1548 | 0.59458 | 3.2006 | 8.078 | 5.0031 | 1.3975 | 7.4742 | 5.4424 | 51.74% |  |  |  | Axis intercept |  | 39.1869 | [17 empty] | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | Hapl. G-copies |  |  | Cq | hapl. gen-copies
| Curve parameters | [28 empty] | R2: |  | 0.99778 | [17 empty] | 24.02 | 48790.1816 |  |  | 34.92 | 34.1828 |  |  | 24.01 | 50470.2653 |  |  | 34.05 | 53.8904
|  |  | R2 ≥ 0.98 |  | Slope: - 3.6 – - 3.1 |  | Accept. criterion | [16 empty] | LOQabs |  | 20 | [23 empty] | 24.05 | 47815.6169 |  |  | 34.42 | 47.668 |  |  | 24.03 | 49815.0964 |  |  | 34.09 | 52.4731
|  |  | Plate B | Plate C | Plate B | Plate C | [43 empty] | 24.06 | 47495.1072 |  |  | 34.63 | 41.4544 |  |  | 24.07 | 48530.1631 |  |  | 34.24 | 47.482
|  | Lectin | 0.99975 | 0.9994 | -3.4236 | -3.5245 | fulfilled | [42 empty] | 23.95 | 51142.1132 |  |  | 34.43 | 47.3521 |  |  | 24.04 | 49490.7083 |  |  | 34.19 | 49.0906
|  | DAS-81419-2 | 0.99726 | 0.99893 | -3.4621 | -3.4557 | fulfilled | [42 empty] | % GMO1 |  |  | 0.08751% |  | sd1 | 0.00013311 |  | %GMO4 |  |  | 0.1024% |  | sd3 | 6.1967e-05
| Trueness | [48 empty] | Extraction 2 |  |  |  | Sample No. | D150167
|  |  | Value [Nominal] | ± 25% | Value [Actual] | Accept. criterion |  | Dynamic range | [41 empty] | Plate B | [7 empty] | Plate C
[8 empty] | 8000 - 20 | [41 empty] | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2 |  | Ref. gene |  | Lectin |  | GM target |  | DAS-81419-2
|  | GM level 1 | 1% | 0.25% | 1.018% | fulfilled |  | fulfilled | [41 empty] | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies |  |  | Cq | hapl. gen-copies
|  | GM level 2 | 0.1% | 0.025% | 0.09292% | fulfilled |  | fulfilled | [41 empty] | 23.91 | 52536.6324 |  |  | 34.59 | 42.5721 |  |  | 23.9 | 54230.7929 |  |  | 34.42 | 42.1154
[50 empty] | 23.99 | 49784.6097 |  |  | 34.42 | 47.668 |  |  | 23.99 | 51134.0509 |  |  | 34.16 | 50.0818
| Relative Repeatability Standard Deviation (RSDr) | [5 empty] | Complement to specificity (Event) | [42 empty] | 24.01 | 49119.4308 |  |  | 34.55 | 43.7198 |  |  | 23.96 | 52146.1342 |  |  | 33.93 | 58.3764
|  |  | RSDr | Accept. Criterion |  |  | Name |  | S/N | Sample No. | % GMO | Conc. (ng/µl) | Cq values | [36 empty] | 23.99 | 49784.6097 |  |  | 34.34 | 50.2729 |  |  | 23.91 | 53877.6504 |  |  | 34.5 | 39.9292
|  | GM level 1 | 7.187% | fulfilled | [45 empty] | %GMO2 |  |  | 0.09164% |  | sd2 | 7.594e-05 |  | %GMO4 |  |  | 0.09019% |  | sd4 | 0.00016072
|  | GM level 2 | 12.41% | fulfilled
| LOD / LOQ (GM target DNA copy number) | [48 empty] | Mean GMO: |  |  | 0.09292% |  |  | sdGM | 0.00011528 |  |  | RSDr | 12.41%
|  | LODabs | 5
|  | LOQabs | 20
[8 empty] | Mass of haploid genome [pg] |  |  | Zygosity ratio
[7 empty] | cotton | 2.33 |  |  | 1
[7 empty] | potato | 1.8 |  |  | 0.5
[7 empty] | linseed | 0.7 |  |  | 0.33333
[7 empty] | maize | 2.6 |  |  | 0.25
[7 empty] | canola | 1.23 |  |  | 0.2
[7 empty] | soybean | 1.15 |  |  | 0.16667
[7 empty] | tomato | 0.99 |  |  | 0.14286
[7 empty] | sugar beet | 1.25 |  |  | 0.125
xx
xx
xx
xx
xx
xx
xx
xx
xx
xx
xx
xx
xx

## Dilution series A
AS-No.: | 150086
Date: | 2015-06-10
Dilution levels for LOD/LOQ determination on plate A
| S/N | #311/1
| Name | DAS-81419-2 10%
| % GMO | 10%
| Ref. gene: | Lectin
Event: |  | DAS-81419-2
| Sample No standard: | D150150 | 5 | Template volume (µl)
|  |  | 123.2 | DNA concentration (ng/µl)
| Species: | soybean | 1.15 | Mass of haploid genome (pg/hapl. genome)
|  |  | 21.4261 | Dilution factor
| Dilution level | GM target DNA copy number | Hapl. genome copy number Reference gene | [7 empty] | 0.2xTE (µl) | Wells/dilution level | Total volume (µl) incl. 20% pipetting reserve | Remaining volume (µl) | Concentration after dilution (ng/µl)
| A | 2500 | 25000 | cp/ | 5 | µl | 1.4468 | µl from | 123,2 ng/µl | + | 29.5532 | 3 | 31 | 20.6 | 5.75
| B | 500 | 5000 | cp/ | 5 | µl | 10.4 | µl from | A | + | 41.6 | 3 | 52 | 24.2
| C | 100 | 1000 | cp/ | 5 | µl | 27.8 | µl from | B | + | 111.2 | 3 | 139 | 38.2
| D | 60 | 600 | cp/ | 5 | µl | 100.8 | µl from | C | + | 67.2 | 10 | 168 | 78.6667
| E | 40 | 400 | cp/ | 5 | µl | 89.3333 | µl from | D | + | 44.6667 | 10 | 134 | 72.5
| F | 20 | 200 | cp/ | 5 | µl | 61.5 | µl from | E | + | 61.5 | 10 | 123 | 70.5
| G | 10 | 100 | cp/ | 5 | µl | 52.5 | µl from | F | + | 52.5 | 10 | 105 | 67.5
| H | 5 | 50 | cp/ | 5 | µl | 37.5 | µl from | G | + | 37.5 | 10 | 75 | 63
| I | 1 | 10 | cp/ | 5 | µl | 12 | µl from | H | + | 48 | 10 | 60 | 60
[11 empty] | 493.7198 | Total TE (µl)

## Dilution series B
AS-No.: | 150086
Date: | 2015-06-11
Standard curve Plate B
| S/N | #311/1
| Name | DAS-81419-2 10%
| % GMO | 10%
| Ref. gene: | Lectin
Event: |  | DAS-81419-2
| Sample No standard: | D150150 | 5 | Template volume (µl)
|  |  | 123.2 | DNA concentration (ng/µl)
| Species: | soybean | 1.15 | Mass of haploid genome (pg/hapl. genome)
|  |  | 6.6957 | Dilution factor
| Dilution level | Hapl. genome copy number Event | Hapl. genome copy number Reference gene | [7 empty] | 0.2xTE (µl) | Wells/dilution level | Total volume (µl) incl. 20% pipetting reserve | Remaining volume (µl) | Concentration after dilution (ng/µl)
| A | 8000 | 80000 | cp/ | 5 | µl | 7.3182 | µl from | 123,2 ng/µl | + | 41.6818 | 6 | 49 | 38.8 | 18.4
| B | 1600 | 16000 | cp/ | 5 | µl | 10.2 | µl from | A | + | 40.8 | 6 | 51 | 39
| C | 400 | 4000 | cp/ | 5 | µl | 12 | µl from | B | + | 36 | 6 | 48 | 38.6
| D | 80 | 800 | cp/ | 5 | µl | 9.4 | µl from | C | + | 37.6 | 6 | 47 | 38
| E | 20 | 200 | cp/ | 5 | µl | 9 | µl from | D | + | 27 | 6 | 36 | 36
[11 empty] | 183.0818 | Total TE (µl)

## Dilution series C
AS-No.: | 150086
Date: | 2015-06-12
Standard curve Plate C
| S/N | #311/1
| Name | DAS-81419-2 10%
| % GMO | 10%
| Ref. gene: | Lectin
Event: |  | DAS-81419-2
| Sample No standard: | D150150 | 5 | Template volume (µl)
|  |  | 123.2 | DNA concentration (ng/µl)
| Species: | soybean | 1.15 | Mass of haploid genome (pg/hapl. genome)
|  |  | 6.6957 | Dilution factor
| Dilution level | Hapl. genome copy number Event | Hapl. genome copy number Reference gene | [7 empty] | 0.2xTE (µl) | Wells/dilution level | Total volume (µl) incl. 20% pipetting reserve | Remaining volume (µl) | Concentration after dilution (ng/µl)
| A | 8000 | 80000 | cp/ | 5 | µl | 7.3182 | µl from | 123,2 ng/µl | + | 41.6818 | 6 | 49 | 38.8 | 18.4
| B | 1600 | 16000 | cp/ | 5 | µl | 10.2 | µl from | A | + | 40.8 | 6 | 51 | 39
| C | 400 | 4000 | cp/ | 5 | µl | 12 | µl from | B | + | 36 | 6 | 48 | 38.6
| D | 80 | 800 | cp/ | 5 | µl | 9.4 | µl from | C | + | 37.6 | 6 | 47 | 38
| E | 20 | 200 | cp/ | 5 | µl | 9 | µl from | D | + | 27 | 6 | 36 | 36
[11 empty] | 183.0818 | Total TE (µl)

## Dilution of samples
AS-No.: | 150086
Date: | 2015-06-11
Dilution of samples with GM level 1 and GM level 2
| Samples |  | Sample No. | DNA conc. of the sample (ng/µl) | Dilution incl. pipetting reserve (4 React.) |  |  | DNA conc. of the dilution (ng/µl) | Required copy number of reference gene / PCR reaction | Mass of haploid genome (pg)
[5 empty] | DNA (µl) |  | 0.2xTE (µl)
| GM level 1 / Extract. 1 |  | D150164 | 106.9 | 10.7577 | + | 89.2423 | 11.5 | 50000 | 1.15 | soybean
| GM level 1 / Extract. 2 |  | D150168 | 142.6 | 8.0645 | + | 91.9355 | 11.5 | 50000 | 1.15 | soybean
| GM level 2 / Extract. 1 |  | D150160 | 110.2 | 10.4356 | + | 89.5644 | 11.5 | 50000 | 1.15 | soybean
| GM level 2 / Extract. 2 |  | D150167 | 132.2 | 8.6989 | + | 91.3011 | 11.5 | 50000 | 1.15 | soybean
Template vol. (µl) |  | 5
Dilution of the samples for the test for specificity:
Name |  | Sample No. | % GMO | DNA conc. of the sample (ng/µl) | Dilution incl. pipetting reserve (1 React.) |  |  | DNA conc. of the dilution (ng/µl) | Required copy number of GM target gene / PCR reaction (target DNA copies) | Mass of haploid genome (pg) |  | Target seq./hapl. genome copies
[5 empty] | Sample (µl) |  | 0.2xTE (µl) |  |  |  | Species
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
|  |  |  | 0 |  | + |  | #N/A |  | #N/A
Template vol. (µl) |  | 5
Zygosity ratio
| cotton |  |  | 1
| potato |  |  | 0.5
| linseed |  |  | 0.33333
| maize |  |  | 0.25
| canola |  |  | 0.2
| soybean |  |  | 0.16667
| tomato |  |  | 0.14286
| sugar beet |  |  | 0.125
| xx |  |  | 0
| xx |  |  | 0
| xx |  |  | 0
| xx |  |  | 0
xx
xx
xx
xx
xx
xx
xx
xx
xx

## Plate layout A “LOD_LOQ”
AS-No.: | 150086
Date: | 2015-06-10
Plate A: Plate layout and preparation of the reaction mix | [7 empty] | Instrument
[6 empty] | Dilution level | Target DNA copies | Mx3005
Cycling programme
[8 empty] | Step | e.g. | T°C | Time (sec) | Acquisition | Cycles
[8 empty] | 1 | UNG
[9 empty] | Denaturation | 95 | 600 | no | 1
[6 empty] | A | 2500 | 2 | Denaturation | 95 | 15 | no | 45
[6 empty] | B | 500 | 3 | Annealing (resp. anneal. + elong.) | 60 | 60 | yes
[6 empty] | C | 100 | 4 | Elongation
[6 empty] | D | 60
[6 empty] | E | 40
[6 empty] | F | 20
[6 empty] | G | 10
[6 empty] | H | 5
[6 empty] | I | 1
Preparation of 100 µl of working solution
for | DAS-81419-2
| Final concentration (nM) | Conc. stock solution (µM) | Conc. working solution (µM) | Stock solution (µl) | Water (µl)
R363 | 400 | 100 | 40 | 40 | 60
R364 | 400 | 100 | 40 | 40 | 60
R365 | 120 | 100 | 12 | 12 | 88
Preparation of reaction mix
| DAS-81419-2 |  | Total consider pipetting reserve!
| Component | µl/reaction | 76 | Date of prep. of working sol. | Lot number
PCR Reaction kit | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | 12.5 | 950 |  | 1412030
Primer F | R363 | 0.25 | 19 | 2015-06-10
Primer R | R364 | 0.25 | 19 | 2015-06-10
Probe | R365 | 0.25 | 19 | 2015-06-10
| H2O | 6.75 | 513
| Total | 20 | 1520
| Template DNA | 5 | µl
| Total reaction volume | 25 | µl
EBC

## Plate layout B “Trueness”
AS-No.: |  | 150086 | [12 empty] | Preparation of 100 µl of working solution
Date: |  | 2016-06-11 | [12 empty] | for | Lectin
Plate B: Plate layout and preparation of the reaction mixes | [15 empty] | Final concentration (nM) | Conc. stock solution (µM) | Conc. working sol. (µM) | Stock solution (µl) | Water (µl)
|  | Standard curve Plate B |  |  |  |  | Samples |  |  |  |  | EBC / NTC |  |  | R010 | 650 | 100 | 65 | 65 | 35
|  | Level A | Level B | Level C | Level D | Level E | GM level 1 Extraction 1 | GM level 1 Extraction 2 | GM level 2 Extraction 1 | GM level 2 Extraction 2 |  |  |  |  | R011 | 650 | 100 | 65 | 65 | 35
[15 empty] | R012B | 180 | 100 | 18 | 18 | 82
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |  | Preparation of reaction mix
Reference gene | A | [14 empty] | Lectin |  | Total consider pipetting reserve! | bar: 1 to 12
B | bar: 1 to 12
[16 empty] | Component | µl/reaction | 40 | Date of prep. of working sol. | Lot No.
C | bar: 1 to 12
[16 empty] | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | 12.5 | 500 |  | 1412030
| D | [13 empty] | PCR Reaction Kit | bar: 1 to 12
[15 empty] | Primer F | R010 | 0.25 | 10 | 2015-03-02
Event | E | [13 empty] | Primer R | R011 | 0.25 | 10 | 2015-03-02 | bar: 1 to 12
[15 empty] | Probe | R012B | 0.25 | 10 | 2015-03-02 | bar: 11
| F | [14 empty] | H2O | 6.75 | 270 | bar: 1 to 12
| C | [14 empty] | Total | 20 | 800 | bar: 11
G | bar: 1 to 12
| E | [14 empty] | Template DNA | 5 | µl | bar: 6 to 9; 11
| H | x | x | x | x | x |  |  |  |  | x | NTC Ref | NTC GM Target |  |  | Total reaction volume | 25 | µl
Instrument |  | 1: Mx3005
Preparation of 100 µl of working solution
|  | 1: Cycling programme |  | for | DAS-81419-2
|  | 1: Step; 3: e.g.; 7: T°C; 9: Time (sec); 11: Acquisition | Cycles |  | Final concentration (nM) | Conc. stock solution (µM) | Conc. working sol. (µM) | Stock solution (µl) | Water (µl)
|  | 1: 1; 3: UNG |  | R363 | 400 | 100 | 40 | 40 | 60
|  | 3: Denaturation; 7: 95; 9: 600; 11: no | 1 | R364 | 400 | 100 | 40 | 40 | 60
|  | 1: 2; 3: Denaturation; 7: 95; 9: 15; 11: no | 45 | R365 | 120 | 100 | 12 | 12 | 88
|  | 3 |  | Annealing (resp. annealing and elongation) |  |  |  | 60 |  | 60 |  | yes
Preparation of reaction mix
|  | 1: 4; 3: Elongation |  |  | DAS-81419-2 |  | Total consider pipetting reserve!
[16 empty] | Component | µl/reaction | 40 | Date of prep. of working sol. | Lot No.
[15 empty] | PCR Reaction Kit | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | 12.5 | 500
[15 empty] | Primer F | R363 | 0.25 | 10 | 2015-06-10
[15 empty] | Primer R | R364 | 0.25 | 10 | 2015-06-10
[15 empty] | Probe | R365 | 0.25 | 10 | 2015-06-10
[16 empty] | H2O | 6.75 | 270
[16 empty] | Total | 20 | 800
[16 empty] | Template DNA | 5 | µl
[16 empty] | Total reaction volume | 25 | µl

## Plate layout C “Trueness”
AS-No.: |  | 150086 | [12 empty] | Preparation of 100 µl of working solution
Date: |  | 2015-06-12 | [12 empty] | for | Lectin
Plate C: Plate layout and preparation of the reaction mixes | [15 empty] | Final concentration (nM) | Conc. stock solution (µM) | Conc. working sol. (µM) | Stock solution (µl) | Water (µl)
|  | Standard curve Plate C |  |  |  |  | Samples |  |  |  | Specificity |  |  |  | R010 | 650 | 100 | 65 | 65 | 35
|  | Level A | Level B | Level C | Level D | Level E | GM level 1 Extraction 1 | GM level 1 Extraction 2 | GM level 2 Extraction 1 | GM level 2 Extraktion 2 |  |  |  |  | R011 | 650 | 100 | 65 | 65 | 35
[15 empty] | R012B | 180 | 100 | 18 | 18 | 82
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |  | Preparation of reaction mix
Reference gene | A | [14 empty] | Lectin |  | Total consider pipetting reserve!
B
[16 empty] | Component | µl/reaction | 40 | Date of prep. of working sol. | Lot No.
C
[16 empty] | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | 12.5 | 500 |  | 1502032
| D | [13 empty] | PCR Reaction kit | bar: 1 to 12
[15 empty] | Primer F | R010 | 0.25 | 10 | 2015-03-02
Event | E | [13 empty] | Primer R | R011 | 0.25 | 10 | 2015-03-02
[15 empty] | Probe | R012B | 0.25 | 10 | 2015-03-02
| F | [14 empty] | H2O | 6.75 | 270
| C | [14 empty] | Total | 20 | 800
G
| E | [14 empty] | Template DNA | 5 | µl | bar: 6 to 9
| H | x | x | x | x | x |  |  |  |  | x | NTC Ref | NTC GM Target |  |  | Total reaction volume | 25 | µl
Instrument |  | 1: Mx3005
Preparation of 100 µl of working solution
|  | 1: Cycling programme |  | for | DAS-81419-2
|  | 1: Step; 3: e.g.; 7: T°C; 9: Time (sec); 11: Acquisition | Cycles |  | Final concentration (nM) | Conc. stock solution (µM) | Conc. working sol. (µM) | Stock solution (µl) | Water (µl)
|  | 1: 1; 3: UNG |  | R363 | 400 | 100 | 40 | 40 | 60
|  | 3: Denaturation; 7: 95; 9: 600; 11: no | 1 | R364 | 400 | 100 | 40 | 40 | 60
|  | 1: 2; 3: Denaturation; 7: 95; 9: 15; 11: no | 45 | R365 | 120 | 100 | 12 | 12 | 88
|  | 3 |  | Annealing (resp. annealing and elongation) |  |  |  | 60 |  | 60 |  | yes
Preparation of reaction mix
|  | 1: 4; 3: Elongation |  |  | DAS-81419-2 |  | Total consider pipetting reserve!
[16 empty] | Component | µl/reaction | 40 | Date of prep. of working sol. | Lot No.
[15 empty] | PCR Reaction kit | TaqMan Universal PCR Master Mix, no AmpErase UNG, #4324018 | 12.5 | 500
[15 empty] | Primer F | R363 | 0.25 | 10 | 2015-06-10
[15 empty] | Primer R | R364 | 0.25 | 10 | 2015-06-10
[15 empty] | Probe | R365 | 0.25 | 10 | 2015-06-10
[16 empty] | H2O | 6.75 | 270
[16 empty] | Total | 20 | 800
[16 empty] | Template DNA | 5 | µl
[16 empty] | Total reaction volume | 25 | µl

## Comments
150086